21st Edition BOD Changes

The 21st Edition of Standard Methods was approved by HQ EPA for Clean Water Act monitoring in early 2007.  This paper points out some of the more significant changes in the order they appear in the method.  Comments are those of the author.

Paragraph (¶) 3g(2) adds allylthiourea (ATU) as an approved nitrification inhibitor.

¶ 4b(4) and 5b allow sample temperature prior to dilution to be 20 ± 3°C rather than ±1°C (but the incubation temperature remains 20 ± 1°C).  Also note that the “prior to dilution” is new.

¶ 4b(5) adds pretreatment procedures for samples containing hydrogen peroxide.

¶ 4c expands guidance on selection and storage of source water and pinpoints the maximum allowed blank depletion at 0.20 mg/L rather the 0.2 mg/L.

¶ 5a expands guidance on preparation of dilution water and recognizes the fact that source water problems can lead to blanks exceeding 0.20 mg/L depletion, even in the absence of any “contamination” in the test process.

¶ 5d adds the following in referring to the amount of seed to add to seeded bottles:  “…if 1 mL of seed…is required to achieve 198 ± 30.5 mg/L BOD in the glucose/glutamic acid check, then use 1 mL in each [seeded]BOD bottle…” (Comment:  Reading the entire subparagraph, one would conclude that if 6 mL of seed were used in the GGA bottle, giving a depletion of 4 mg/L DO, the same amount would be used in each seeded bottle, also causing a depletion of 4 mg/L.  This means the seed and not the sample would be by far the major contributor to DO depletion in seeded samples with low BOD, such as effluents.  Because of imprecision inherent to the BOD test, the seed contribution to DO depletion could completely mask the sample contribution for low BOD samples.  This seems contrary to good reason.  A better procedure is to use the volume of seed in the seed control bottle that gives 2.0 or more mg/L depletion making it a “valid” bottle, and in seeded bottles, use the volume of seed that will produce a seed correction of 0.6 – 1.0 mg/L, just as it says to do in the method.)

¶ 5i allows a ± 6-hour variance in the “5-day” incubation period (i.e., it is now 5 days ± 6 hours).

¶ 6b adds “the glucose/glutamic acid check is the primary basis for establishing accuracy and precision of the BOD test…”  (Comment:  Paragraph 8 confuses this issue by retaining the old statement that “there is no measurement for establishing bias of the BOD procedure.”  Since “accuracy” as used in 6b contains a “bias” and “precision” component, it appears that paragraph 6b recognizes that the 198 mg/L objective for the GGA test is a bias objective.)

¶ 7b establishes the following circumstances under which results should be “identified” in reports:

- Blank exceeds 0.20 mg/L;

- GGA result is outside acceptance limits;

- Test replicates show more that 30% difference between high and low values (Comment:  elsewhere, the term “dilutions” is used rather than “replicates” and it is probably “dilutions” that is intended although the two terms are not synonymous.);

- Seed control samples do not deplete at least 2.0 mg/L, with a retention of at least 1.0 mg/L DO, or

- The minimum DO retained is less than 1.0 mg/L.

(Comment:  The method does not say what is meant by “identifying” the results, but it probably means mentioning the result(s) in the remarks section of a report.  Also, the method does not say results must be “identified” if no dilution depletes at least 2.0 mg/L.  This probably means that reporting such results with the “<” less than symbol is sufficient.)

¶ 8b establishes a lower detection limit of 0 mg/L, replacing the 2.0 mg/L lower limit in previous editions.  The 0 mg/L is possible when all the DO depletion in a seeded, undiluted bottle is contributed by the seed.

There are other, less significant changes in the 21st Edition and, because many of the paragraphs were restructured, it would be difficult to assure that all significant changes are mentioned here.  The reader is encouraged to study the new method in detail.